## Slide 1: BPMLHC
BPMLHC = BPMI (same electronics + FIFO)
Acquisition modes
  FIFO (work-horse for Injection Test)
  In parallel/parasitically:
    Synchronous capture
    (XPOC – post-mortem)
Good news and Issues
Outlook for future tests

## Slide 2: FIFO mode
Asynchronous FIFO (no fine timing needed)
  Triggered with timing events (as for LTI)
Start (-100 ms)
Stop/Read (+10 ms)
(don’t read too early!)
BEAM
Read FIFO values (max = 1000), calibrate, calculate average position -> publish (LSA concentrator) ->

## Slide 3: Capture mode
Synchronous (fine BST synchronization needed)
  Only acquisition mode that allows bunch-selection (injected vs. circulating beam)
  Requires selecting bunches and turns (operational settings)
  Requires setting-up of bunch-delays and 40MHz phases
  At injection will be triggered using LHC injection pre-pulses over BST provided by AB-RF
    Produced sufficiently early (and without affecting the AB-BT pre-pulses) ..
  Everywhere else in LHC cycle it will be triggered with LHC timing event
    Issue with LSA concentrator to be solved ?

## Slide 4: Capture results (B1RA system)
LHC Position | Position | ADC-Data | TDC | Bunch-ID
BPMSW.1R2.B1 | 32507 | 689 | 4 | 450
BPMSX.4R2.B1 | 32506 | 0 | 7 | 449
BPMYB.4R2.B1 | 32538 | 530 | 6 | 481
BPM.6R2.B1 | 32568 | 466 | 3 | 511
BPM.8R2.B1 | 32576 | 0 | 0 | 519
BPM.10R2.B1 | 32603 | 417 | 6 | 546
BPM.12R2.B1 | 32638 | 512 | 5 | 581
BPM.14R2.B1 | 32673 | 0 | 1 | 616
BPM.16R2.B1 | 32709 | 526 | 5 | 652
BPM.18R2.B1 | 32744 | 0 | 1 | 687
BPM.20R2.B1 | 32780 | 523 | 4 | 723
BPM.22R2.B1 | 32815 | 355 | 3 | 758
BPM.24R2.B1 | 32850 | 0 | 0 | 793
BPM.26R2.B1 | 32886 | 631 | 4 | 829
BPM.28R2.B1 | 32921 | 492 | 2 | 864
BPM.30R2.B1 | 32957 | 0 | 7 | 900
BPM.32R2.B1 | 32992 | 663 | 4 | 935
BPM.34R2.B1 | 33027 | 548 | 5 | 970
To get valid capture for TDC (1>X>7) requires changing 40MHz phase
Realignment of bunches requires changing bunch delay
These ‘expert’ settings that will hopefully not change very often
Algorithms presently being thought about

## Slide 5: XPOC (PM) mode
Mainly used in LSS6 for LBDS systems:
  For XPOC analysis, to observe closed-orbit position (turn-by-turn) prior to an emergency beam dump (BPMSE ++)
  Basically seems to work fine, pending issue with alignment of turns (1024) in low-level software

## Slide 6: Good news & Issues
BPM calibration factors are OK to better than 10%
Workable intensity range as expected (low limit = 1.5E9 charges/bunch).
  Low intensity pilots can be used for commissioning
Also important for physics beams (protons/ions).
  Will involve a change of sensitivity below 4E10 charges/bunch (average)
  Noise figures to be measured with beam
BPM Issues:
BPMI. 29504.V didn’t trigger
BPMSW.1L2.V noisy (was disabled)
BPM.33R2.V with H/V cable inversion
Wrong geometrical correction for warm BPMs (pt3), now OK

## Slide 7: BPM Intensity modules
Were installed on Q8R2 and Q10L3
Less sensitive than expected due to non-linearity
  These effects were not seen on previous hardware version (TT40/TT60/TI2)
Will be corrected (polynomial) to provide errors below 5%
  A bit of commissioning with beam needed to get parameters right.
Chart: High-Sensitivity Transfer Function for BPM Intensity module
| 1.0808358315340907 | 1.2113223595514648 | 1.3575621902418851 | 1.5214571792902174 | 1.705138788523048 | 1.9109958056671967 | 2.1417054106433273 | 2.4002679924132271 | 2.6915951300116001 | 3.0182814447391637 | 3.4021991665317306 | 3.8349502459170046 | 4.2979339184686705 | 4.8168124181507892 | 5.3983337835771801 | 6.0500607266949595 | 6.7804689861992635 | 7.5990575549024628 | 8.5952741617467456 | 9.7220921649591983 | 10.89581532869799 | 12.211239068991931 | 13.685470531730607 | 15.337682164495373 | 17.19925889530186 | 19.286780321500583 | 21.740009716211343 | 24.50528364934431 | 27.463743471078111 | 30.779370532418561 | 34.495284714903626 | 38.659811652384775 | 43.337088792811585
 | 8492 | 8503.5 | 8513 | 8521.9 | 8532.8 | 8544.4 | 8559.7 | 8571.9 | 8589.1 | 8607.1 | 8635.7 | 8671.6 | 8722.3 | 8792 | 8886.1 | 8986.3 | 9075.2 | 9154.5 | 9231.8 | 9309.8 | 9398.4 | 9480.6 | 9574.3 | 9679.4 | 9808.9 | 9962.4 | 10194.4 | 10403.6 | 10615 | 10793.9 | 11015.7 | 11226.6 | 11503.8
 | 8390 | 11770
 | 8390 | 11770

## Slide 8: Outlook coming injection tests
BPM hardware commissioning on-going:
  LSS8, 7-8, LSS7 (6-7, LSS6, TD62)
  Required accesses requested
BPM Intensity modules to be installed:
  Q8R7, Q10L7 and on BPMD (in TD62)
Dry-run to be planned next week